## Slide 1: Assay Allocation Algorithm Explainer and Status
Pete Howard
Thursday 16th Nov 2017

## Slide 2: The N+3 Concept
In other words, why do I claim that we must allocate 6 of each assay, if we want to cope with 3 targets being present at once…

## Slide 3: The N+3 Concept
1: A B | 2: A C | 3: A D | 4: A E | 5: A F | 6: A G
Context
We are trying to detect the presence of target A
The assay for A is allocated to chambers: {1,2,3,4,5,6}
We admit that up to 3 targets may be present at the same time
I.e. N=3
We have deployed the assay for A to 6 chambers
I.e. N+3
How many fired? | Interpretation
6:6 | A is present. No malfunctions.
5:6 | A is present. But one of {1,2,3,4,5,6} failed to fire, or was called wrong.
4:6 | A is not present (see 3:6). But one of {1,2,3,4,5,6} fired when it shouldn’t have, or was called wrong.
3:6 | A is not present. No malfunctions. 3 of {1,2,3,4,5,6) fired due to the presence of N other targets. {B,C,D} for example.

## Slide 4: Will N+2 Suffice? (Slide 1 of 2)
1: A B | 2: A C | 3: A D | 4: A E | 5: A F
Context
We are trying to detect the presence of target A
The assay for A is allocated to chambers: {1,2,3,4,5}
We admit that up to 3 targets may be present at the same time
I.e. N=3
We have deployed the assay for A to 5 chambers
I.e. N+2
How many fired? | Interpretation
5:5 | A is present. No malfunctions.
4:5 | Either A is present. But one of {1,2,3,4,5} failed to fire, or was called wrong. Or: A is not present. {1,2,3} (for example), fired due to the presence of N other targets. E.g. {B,C,D}, But also one of {4,5} fired when it shouldn’t have, or was called wrong.
3:5 | A is not present. No malfunctions. 3 of {1,2,3,4,5) fired due to the presence of N other targets. {B,C,D} for example.

## Slide 5: Will N+2 Suffice? (Slide 2 of 2)
Is the probability of a failure to fire, very different from the probability of a spurious fire?
The probability of the “Or” case includes two probabilities multiplied – which lowers it.
If so could make a call based on the higher probability
Can we avoid the first case by exploiting the DAC test?
4:5 | Either A is present. But one of {1,2,3,4,5} failed to fire, or was called wrong. Or: A is not present. {1,2,3} (for example), fired due to the presence of N other targets. E.g. {B,C,D}, But also one of {4,5} fired when it shouldn’t have, or was called wrong.

## Slide 6: What Limits our Capacity? (1:4)
Example:
I have 4 assays, and 10 chambers.
I want to cope with up to N=3
I.e. I must deploy 6 of each assay.
I choose to deploy A first.
I choose a set of 6 chambers – e.g. {2,4,5,6,7,10}
That choice was one of 210 choices available.
See this handy “n choose k” website to see where 210 comes from.
https://www.hackmath.net/en/calculator/n-choose-k

## Slide 7: What Limits our Capacity? (2:4)
What happens when I move on to choose 6 chambers for B?
Some of the 209 remaining possible choices must be ruled out because we chose {2,4,5,6,7,10} for A.
For example this set must be ruled out: {1,2,3,4,8,10}
Note the overlap (intersection) is just {2,4,10}
Why must it be ruled out..?
Next page.

## Slide 8: What Limits our Capacity? (3:4)
Given {2,4,5,6,7,10} for A, why must {1,2,3,4,8,10} be ruled out for B?
1: B | 2: A B | 3: B | 4: A B | 5: A C
6: A D | 7: A | 8: B | 9: B | 10: A B
The presence of {B,C,D} would cause a false positive for A – as shown by this example:
B causes {2,4,10} to fire.
C causes 5 to fire.
D causes 6 to fire.
This is: 5:6 of A’s chambers, which we previously characterised as a positive call for A.

## Slide 9: What Limits our Capacity? (4:4)
The root cause is that the chamber sets intersect by 3 (N) members: {2,4,10}
Had the overlap been 2 (N-1), the clash cannot happen.
Remedy for the B-after-A case
After A has been allocated, we must purge from our pool of choices for B, any set in the pool that intersects with that used for A, by more than N-1.
Remedy extrapolated to all assay allocations
After assay (xxx) has been allocated, we must purge from our pool, any set in the pool that intersects with that used for any assay allocated prior to xxx, by more than N-1.
Ps. This is more or less the how the current software tool works, except that it lags behind the thinking in this document a bit. Because it preceded the malfunction resilience thinking. It only allocates N+1, and purges intersections larger than 1. (Very easy to upgrade).

## Slide 10: Modelling how far a given sized pool gets us (1:2)
We have two possible approaches to characterising how far we can get for a given experiment parameters.
  Algebraic
  Computational trial and error to see how far we get with different configurations before it fails.
The current software does (2). – but:
  needs to be updated a bit to catch up with a few points mentioned in this PPT
  has problematic runtimes for large enough numbers.
If we can crack (1), we get instant (to all intents and purposes) answers to model the entire trade off space.
We have been struggling with (1) because of the limit of our mathematics capability.
Cont…

## Slide 11: Modelling how far a given sized pool gets us (2:2)
Arthur has contributed (yesterday) a new suggested heuristic for making the allocations, and a new way to formulate the maths. (A kind of sequential recipe for doing the allocations).
It doesn’t change any of the fundamentals and facts covered in the PPT. It’s like a new perspective on the same problem.
But the combination of the heuristic and the new maths formulation WILL, I think lead us now quite quickly to the algebraic model that will reveal everything.
In the meantime, I have briefed an expert mathematician Chris Tofts, and he’s coming to help us Monday.
And unless I get this math model working fairly quickly, I’ll fix up the existing software and set it running to at least begin to give us some trade off design points to consider.